# FEV2 front-end board – bug report

That application note reports detected bugs on FEV2 front-end board (production # 0321) and details how it will be corrected in FEV2_2 front-end board (production number # 0408)

- Floating VDDA : VDDA not connected to VDD

  - To be connected to VDD

- Calibration differential voltage to single ended current is not working : Voltage driven current source of calibration circuit is not working due to a wrong polarization

  - Schematic to be modified. Differential voltage will first be converted to single ended voltage then to a current.

- Calibration crosstalk between Ch.1 and Ch. 5 is abnormally high : crosstalk between Ch.1 and Ch.5 is abnormally high due to mutual induction of associated inductors.

  - Placement of inductors to be modified to ensure better EMC.

- Calibration crosstalk between Ch.1 and Ch.2 is abnormally high : crosstalk between Ch.1 and Ch.2 is abnormally high due to line-to-pad parasitic capacitance.

  - Routing of calibration lines to be modified to ensure better EMC.

- Calibration linearity slope is too low : calibration linearity slope is around 6mV/mV instead of 12mV/mV theoretic due to a too low inductor value.

  - Calibration inductors value to be modified from 10uH to 1mH

Please report any bug to fleury@lal.in2p3.fr